## 표지
Park NX-Wafer
END USER: Huahong Grace Fab3 #2
SID No. : C160112-020725

## 통합_(NX-Wafer 200mm)
# | Module | Check Items | Min | Criteria | Max | Measurement | Unit | PASS/FAIL | Category | Trend | 공통/옵션 | DB_Key | Remark | 중간검수
1 | Performance | Specification of Step Height Sample for repeatability test |  | - |  | 177.5±2.0 | nm | Pass | Performance |  |  |  | 아래 측정(AFM,AFP Repeatability) 에 사용하는 Sample(180nm) spec 기입 ex)183.8 ± 2.0nm
2 | Performance | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) |  |  | 1 | 0.36 | nm | Pass | Performance | Check item |  |  | 측정시 Sample EFEM으로 로딩하는 Job으로 진행 필요 3.13.0 부터 해당레시피 변경됨
3 | Performance | AFM, Static Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) |  |  | 0.5 | 0.07 | % | Pass | Performance | Check item |  |  | 측정시 Sample EFEM으로 로딩하는 Job으로 진행 필요 3.13.0 부터 해당레시피 변경됨
4 | Performance | Cycle time for 15-point measurement |  | - |  | 1773.939 | sec | Pass | Performance | Check item |  |  | Static Repeatability w/ 180nm Step Height Sample 측정 Data log로 확인, 가이드 참고하여 분석시트 사용, Mask, Panel는 N/A 처리
5 | Performance | Specification of Step Height Sample for short-term stability test |  | - |  | 452.5±2.7 | nm | Pass | Performance |  |  |  | 아래 측정(Short-term stability) 에 사용하는 Sample spec 기입 ex)450 ± 2.0nm, wafer sample 사용(180nm or 450nm)
6 | Performance | Short-term stability using Step Height Sample (3σ) |  |  | 1 | 0.3 | nm | Pass | Performance | Check item |  |  | Cal 450nm, 450nm or 180nm의 Sample을 활용하여 72시간(3days) 측정 진행 1recipe(100 repeats)를 21번 반복 측정 - 총 2100 point의 3시그마 사용, Guide 및 분석시트 사용
7 | Performance | Roughness Repeatability Test 1x1um(10Repeats) |  | - |  | - | - | Pass | Performance |  |  |  | NX-Mask는 해당 테스트 삭제
8 | Performance | Roughness Repeatability 1x1um(RSD%) |  |  | 5 | 1.6 | % | Pass | Performance | Check item |  |  | NX-Mask는 해당 테스트 삭제, 1point에서 10번 측정, 분석시트 사용하여 RSD값 및 1σ 기입
9 | Performance | Roughness Repeatability 1x1um(1σ) |  |  | 0.01 | 0.004 | nm | Pass | Performance | Check item |  |  | NX-Mask는 해당 테스트 삭제, 1point에서 10번 측정, 분석시트 사용하여 RSD값 및 1σ 기입
10 | Performance | Roughness Repeatability Test 80x80um(10Repeats) |  | - |  | - | - | Pass | Performance |  |  |  | NX-Mask는 해당 테스트 삭제
11 | Performance | Roughness Repeatability 80x80um(RSD%) |  |  | 5 | 0.2 | % | Pass | Performance | Check item |  |  | NX-Mask는 해당 테스트 삭제, 1point에서 10번 측정, 분석시트 사용하여 RSD값 및 1σ 기입
12 | Performance | Roughness Repeatability 80x80um(1σ) |  |  | 0.01 | 0.001 | nm | Pass | Performance | Check item |  |  | NX-Mask는 해당 테스트 삭제, 1point에서 10번 측정, 분석시트 사용하여 RSD값 및 1σ 기입
13 | Performance(w/ EFEM) | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, 3σ) |  |  | 1 | 0.29 | nm | Pass | Performance | Check item |  |  | 1day로 변경
14 | Performance(w/ EFEM) | AFM, Dynamic Repeatability w/ 180nm Step Height Sample (15 repeats, RSD%) |  |  | 0.5 | 0.06 | % | Pass | Performance | Check item |  |  | 1day로 변경
15 | SW | Simulated Automation Test |  | Yes |  | Yes | - | Pass | Performance
16 | System Noise | Lateral(C-D) noise |  |  | 0.25 | 0.141 | nm | Pass | Performance | Check item |  |  | B100-NCH Calibration 후 측정, Panel 장비는 Max spec. 삭제
17 | System Noise | Min. Vibration Transmissbility (~100Hz) |  |  | 0.1 | 0.04 | - | Pass | Performance | Check item |  |  | Floor vs AFM Base, Vario사용시 Max spec. 0.2, 그외는 0.1, 진동센서 없는 경우 N/A | ☑
18 | System Noise (NX-Wafer 200mm, w/o Sliding Stage, HDM, PTR) | RMS of ZeroScan |  |  | 0.05 | 0.035 | nm | Pass | Performance | Check item |  |  | Z Drive(Long head), Z Height(standard head) | ☑
19 | AFM Head | Beam intensity with AC160TS cantilever | 2.5 |  | 3.5 | 2.94 | V | Pass | Performance |  |  |  |  | ☑
20 | AFM Head | Check if the driving % changes proportionally as the NCM amplitude changes |  | Yes |  | Yes | - | Pass | Performance |  |  |  | Manual 참고, Amp. 1nm, 20nm, 100nm 일때 확인, 아래 값이 비례적으로 변하는지 확인한 후 Yes 진행 | ☑
21 | AFM Head | NCM driving % at an amplitude of 1 nm |  | - |  | 0.08 | % | Pass | Performance |  |  |  | liquid mode 활성화 확인 필요 및 Amplitude 1nm로 일때의 Driving%, Capture 후 저장 | ☑
22 | AFM Head | NCM driving % at an amplitude of 20 nm | 0.1 |  | 50 | 1.63 | % | Pass | Performance |  |  |  | liquid mode 활성화 확인 필요 및 Amplitude 20nm로 일때의 Driving%, Capture 후 저장 | ☑
23 | AFM Head | NCM driving % at an amplitude of 100 nm |  | - |  | 8.31 | % | Pass | Performance |  |  |  | liquid mode 활성화 확인 필요 및 Amplitude 100nm로 일때의 Driving%, Capture 후 저장 | ☑
24 | AFM Head | Check Z-height non-linearity correction |  | Yes |  | Yes | - | Pass | Performance |  |  |  |  | ☑
25 | AFM Head | Specification of Step Height Sample for Z Linearity by Z scanner position test |  | - |  | 461.2±2.8 | nm | Pass | Performance |  |  |  | 아래 측정(Z Linearity by Z scanner position test) 에 사용하는 Sample(450nm) spec 기입 ex)450 ± 2.0nm | ☑
26 | AFM Head | Z Linearity by Z scanner position |  |  | 2 | 0.37 | % | Pass | Performance | Check item |  |  | Z non-linearity corrction test 값으로 사용, Long head(2%) : Z scannner position -6,0,+6, Standard head(2%) : Z scannner position -3,0,+3 | ☑
27 | AFM Head | Full range of Z Detector during Z Scan sweep | 10 |  | 18 | 13.15 | V | Pass | Performance | Check item |  |  | servo off, TSH 인경우 MIN11.8, MAX15.4 | ☑
28 | AFM Head | Perform AD calibration and verify the sensitivity value | 0.9 | 1 | 1.1 | 1.033 | - | Pass | Performance | Check item |  |  | AC160TS로 MicrometerPerErrorSignalAdc(A-B Sensitivity)가 0.2일때 AD Cal 진행하여 Sensitivity값이 1±0.1안에 드는지 확인 | ☑
29 | AFM Head(Auto Align) | A-B after Auto PSPD Alignment | -0.5 |  | 0.5 | -0.084 | V | Pass | Performance |  |  |  |  | ☑
30 | AFM Head(Long) | Full range of Z Height during Z Scan sweep | 39 |  | 46 | 41.1222 | um | Pass | Performance | Check item |  |  | servo off | ☑
31 | AFM Head(Long) | Z Height noise |  |  | 0.1 | 0.066 | nm | Pass | Performance | Check item |  |  |  | ☑
32 | AFM Head | Z Linearity by 4 different sample height | 0.98 |  | 1.02 | 1.0043 | - | Pass | Performance | Check item |  |  | 4 different Step Height Sample | ☑
33 | AFM Head | - 940nm Step Height Sample (10 Repeats) | 943 | 948.6 | 954.2 | 951.9499 | nm | Pass | Performance |  |  |  | Z scanner Recipe사용, Sample spec Criteria에 기입, 측정 Sample spec 정확히 기입 | ☑
34 | AFM Head | - 200~500nm Step Height Sample (10 Repeats) | 458.4 | 461.2 | 464 | 461.79 | nm | Pass | Performance |  |  |  | Z scanner Recipe사용, Sample spec Criteria에 기입, 측정 Sample spec 정확히 기입 | ☑
35 | AFM Head | - 50~200nm Step Height Sample (10 Repeats) | 181.8 | 183.8 | 185.8 | 183.4169 | nm | Pass | Performance |  |  |  | Z scanner Recipe사용, Sample spec Criteria에 기입, 측정 Sample spec 정확히 기입 | ☑
36 | AFM Head | - Under 50nm Step Height Sample (10 Repeats) | 14.3 | 15.1 | 15.9 | 14.63 | nm | Pass | Performance |  |  |  | Z scanner Recipe사용, Sample spec Criteria에 기입, 측정 Sample spec 정확히 기입 | ☑
37 | Focus Stage | OpticalAxisShiftPerFocusX | -0.01 |  | 0.01 | -0.0013 | - | Pass | Performance | Check item |  |  |  | ☑
38 | Focus Stage | OpticalAxisShiftPerFocusY | -0.01 |  | 0.01 | -0.0007 | - | Pass | Performance | Check item |  |  |  | ☑
39 | Focus Stage | OpticalAxisShiftPerFocusX repeatability (10repeats, 1σ) |  |  | 0.0015 | 0.0001 | - | Pass | Performance | Check item |  |  |  | ☑
40 | Focus Stage | OpticalAxisShiftPerFocusY repeatability (10repeats, 1σ) |  |  | 0.0015 | 0.00015 | - | Pass | Performance | Check item |  |  |  | ☑
41 | Sample Chuck | Check sample surface mapping |  | N/A |  | N/A | um | Pass | Performance | Check item |  |  | Sliding Stage없는 경우 N/A(Overlay제외, (Overlay Spec. 50 µm)), Sim tape 작업 후 진행 (Spec: <30 µm(@300 mm, Anti-warpage chuck) / 그 외 및 Wafer+R-Stage 사용 시 <50 µm)
42 | Vision | Check the Vision brightness calibration |  | Yes |  | Yes | - | Pass | Performance |  |  |  |  | ☑
43 | Vision | LightStrengthGain |  |  |  | 0.23 | - | Pass | Performance | Check item |  |  | Vision brightness cal 후 DB 값 기입 Vision-Camera-LightStrengthGain, Cal 값 기입 후 출고시 1로 설정해서 출하 | ☑
44 | XY Scanner | Full range of X Detector during X Scan sweep | 150 |  | 240 | 185.6 | um | Pass | Performance | Check item |  |  | servo off | ☑
45 | XY Scanner | Full range of X1 Detector during X Scan sweep | 150 |  | 240 | 180.6 | um | Pass | Performance | Check item |  |  | servo off | ☑
46 | XY Scanner | Full range of X2 Detector during X Scan sweep | 150 |  | 240 | 190.6 | um | Pass | Performance | Check item |  |  | servo off | ☑
47 | XY Scanner | Full range of Y Detector during Y Scan sweep | 150 |  | 240 | 186.5 | um | Pass | Performance | Check item |  |  | servo off | ☑
48 | XY Scanner | Full range of Y1 Detector during Y Scan sweep | 150 |  | 240 | 188.7 | um | Pass | Performance | Check item |  |  | servo off | ☑
49 | XY Scanner | Full range of Y2 Detector during Y Scan sweep | 150 |  | 240 | 184.4 | um | Pass | Performance | Check item |  |  | servo off | ☑
50 | XY Scanner | Full range of X1 Detector Raw during X Scan sweep | 11 |  | 17 | 14.48 | V | Pass | Performance | Check item |  |  | servo off | ☑
51 | XY Scanner | Full range of X2 Detector Raw during X Scan sweep | 11 |  | 17 | 14.56 | V | Pass | Performance | Check item |  |  | servo off | ☑
52 | XY Scanner | Full range of Y1 Detector Raw during Y Scan sweep | 11 |  | 17 | 14.66 | V | Pass | Performance | Check item |  |  | servo off | ☑
53 | XY Scanner | Full range of Y2 Detector Raw during Y Scan sweep | 11 |  | 17 | 15.5 | V | Pass | Performance | Check item |  |  | servo off | ☑
54 | XY Scanner | XY Orthogonality | 89.7 | 90 | 90.3 | 90 | degree | Pass | Performance | Check item |  |  | Full stroke image의 ortho 확인 | ☑
55 | XY Scanner | X Non-linearity error | 0 |  | 4 | 2.95 | % | Pass | Performance | Check item |  |  |  | ☑
56 | XY Scanner | Y Non-linearity error | 0 |  | 4 | 2.91 | % | Pass | Performance | Check item |  |  |  | ☑
57 | XY Scanner | X axis OPM |  |  | 4 | 1.06 | nm | Pass | Performance | Check item |  |  | @80x80um | ☑
58 | XY Scanner | Y axis OPM |  |  | 4 | 0.62 | nm | Pass | Performance | Check item |  |  | @80x80um | ☑
59 | XY Scanner | Check the full stroke image(servo on) |  | Yes |  | Yes | - | Pass | Performance |  |  |  | scan size : 100x100um, overscan : 0%, 이미지 측정하여 왜곡 확인 (orthogonality 측정시랑 같은 조건으로 scan size 및 overscan 만 변경) | ☑
60 | XY Scanner | X Ringing(Rpv) |  |  | 1 | 0.36 | nm | Pass | Performance | Check item |  |  | Recipe사용 | ☑
61 | XY Scanner | Y Ringing(Rpv) |  |  | 1 | 0.49 | nm | Pass | Performance | Check item |  |  | Recipe사용 | ☑
62 | XY Scanner | X Detector noise (XY Servo on) |  |  | 0.4 | 0.34 | nm | Pass | Performance | Check item |  |  | Anti-warpage 샘플척은 0.5, 그외는 0.4 | ☑
63 | XY Scanner | Y Detector noise (XY Servo on) |  |  | 0.4 | 0.35 | nm | Pass | Performance | Check item |  |  | Anti-warpage 샘플척은 0.6, 그외는 0.4 | ☑
64 | XY Scanner | X1 Detector noise (XY Servo on) |  |  | 0.5 | 0.46 | nm | Pass | Performance | Check item |  |  | Anti-warpage 샘플척은 0.6, 그외는 0.5 | ☑
65 | XY Scanner | Y1 Detector noise (XY Servo on) |  |  | 0.5 | 0.45 | nm | Pass | Performance | Check item |  |  | Anti-warpage 샘플척은 0.7, 그외는 0.5 | ☑
66 | XY Scanner | X2 Detector noise (XY Servo on) |  |  | 0.5 | 0.45 | nm | Pass | Performance | Check item |  |  | Anti-warpage 샘플척은 0.6, 그외는 0.5 | ☑
67 | XY Scanner | Y2 Detector noise (XY Servo on) |  |  | 0.5 | 0.46 | nm | Pass | Performance | Check item |  |  | Anti-warpage 샘플척은 0.7, 그외는 0.5 | ☑
68 | XY Scanner | ResonanceFrequencyHz |  | 200 |  | 200 | Hz | Pass | Performance | Check item |  |  | 200mm XY Scanner는 200, 300mm XY Scanner는 160 | ☑
69 | XY Scanner | ServoCutoffFrequencyHz |  | 80 |  | 80 | Hz | Pass | Performance | Check item |  |  | 200mm XY Scanner는 80, 300mm XY Scanner는 50 | ☑
70 | XY Stage | Image offset after moving from -X to +X (X scan) |  |  | 200 | 60.79 | nm | Pass | Performance | Check item |  |  | Plus direction move Drift, X move X Scan(XX) 700sec, HDM/PSS는 N/A
71 | XY Stage | Image offset after moving from -X to +X (Y scan) |  |  | 200 | 63.72 | nm | Pass | Performance | Check item |  |  | Plus direction move Drift, X move Y Scan(XY) 700sec, HDM/PSS는 N/A
72 | XY Stage | Image offset after moving from -Y to +Y (Y scan) |  |  | 200 | 32.95 | nm | Pass | Performance | Check item |  |  | Plus direction move Drift, Y move Y Scan(YY) 700sec, HDM/PSS는 N/A
73 | XY Stage | Image offset after moving from -Y to +Y (X scan) |  |  | 200 | 19.04 | nm | Pass | Performance | Check item |  |  | Plus direction move Drift, Y move X Scan(YX) 700sec, HDM/PSS는 N/A
74 | XY Stage | Image offset after moving from +X to -X (X scan) |  |  | 200 | 34.42 | nm | Pass | Performance | Check item |  |  | Minus direction move Drift, X move X Scan(XX) 700sec, HDM/PSS는 N/A
75 | XY Stage | Image offset after moving from +Y to -Y (Y scan) |  |  | 200 | 194.09 | nm | Pass | Performance | Check item |  |  | Minus direction move Drift, Y move Y Scan(YY) 700sec, HDM/PSS는 N/A
76 | XY Stage | Image offset after moving from +X to -X (Y scan) |  |  | 200 | 104 | nm | Pass | Performance | Check item |  |  | Minus direction move Drift, X move Y Scan(XY) 700sec, HDM/PSS는 N/A
77 | XY Stage | Image offset after moving from +Y to -Y (X scan) |  |  | 200 | 21.97 | nm | Pass | Performance | Check item |  |  | Minus direction move Drift, Y move X Scan(YX) 700sec, HDM/PSS는 N/A
78 | XY Stage | XY Stage positioning X offset (Vision Pattern Matching) |  |  | 1 | 0.72 | um | Pass | Performance | Check item |  |  | Recipe/분석프로그램 사용, Metro cal 사용, HDM/PSS/PTR/Mask는 N/A
79 | XY Stage | XY Stage positioning Y offset (Vision Pattern Matching) |  |  | 1 | 0.83 | um | Pass | Performance | Check item |  |  | Recipe/분석프로그램 사용, Metro cal 사용, HDM/PSS/PTR/Mask는 N/A
80 | XY Stage | XY Stage positioning X 1σ (Vision Pattern Matching) |  |  | 0.2 | 0.15 | um | Pass | Performance | Check item |  |  | Recipe/분석프로그램 사용, Metro cal 사용, HDM/PSS/PTR/Mask는 N/A
81 | XY Stage | XY Stage positioning Y 1σ (Vision Pattern Matching) |  |  | 0.2 | 0.17 | um | Pass | Performance | Check item |  |  | Recipe/분석프로그램 사용, Metro cal 사용, HDM/PSS/PTR/Mask는 N/A
82 | XY Stage | XY Stage positioning X offset (In-Die Align) |  |  | 2 | 0.67 | um | Pass | Performance | Check item |  |  | Recipe/분석프로그램 사용, Metro cal 사용, HDM/PSS/PTR/Mask는 N/A
83 | XY Stage | XY Stage positioning Y offset (In-Die Align) |  |  | 2 | 0.78 | um | Pass | Performance | Check item |  |  | Recipe/분석프로그램 사용, Metro cal 사용, HDM/PSS/PTR/Mask는 N/A
84 | XY Stage | XY Stage positioning X 1σ (In-Die Align) |  |  | 0.4 | 0.15 | um | Pass | Performance | Check item |  |  | Recipe/분석프로그램 사용, Metro cal 사용, HDM/PSS/PTR/Mask는 N/A
85 | XY Stage | XY Stage positioning Y 1σ (In-Die Align) |  |  | 0.4 | 0.16 | um | Pass | Performance | Check item |  |  | Recipe/분석프로그램 사용, Metro cal 사용, HDM/PSS/PTR/Mask는 N/A
86 | XY Stage | XY Stage positioning X offset (LLC Translation) |  |  | 4 | 1.43 | um | Pass | Performance | Check item |  |  | Recipe/분석프로그램 사용, Metro cal 사용, HDM/PSS/PTR/Mask는 N/A
87 | XY Stage | XY Stage positioning Y offset (LLC Translation) |  |  | 4 | 0.98 | um | Pass | Performance | Check item |  |  | Recipe/분석프로그램 사용, Metro cal 사용, HDM/PSS/PTR/Mask는 N/A
88 | XY Stage | XY Stage positioning X 1σ (LLC Translation) |  |  | 0.8 | 0.29 | um | Pass | Performance | Check item |  |  | Recipe/분석프로그램 사용, Metro cal 사용, HDM/PSS/PTR/Mask는 N/A
89 | XY Stage | XY Stage positioning Y 1σ (LLC Translation) |  |  | 0.8 | 0.21 | um | Pass | Performance | Check item |  |  | Recipe/분석프로그램 사용, Metro cal 사용, HDM/PSS/PTR/Mask는 N/A
90 | XY Stage | XY Stage positioning X offset (Global Align) |  |  | 6 | 2.55 | um | Pass | Performance | Check item |  |  | Recipe/분석프로그램 사용, Metro cal 사용, HDM/PSS/PTR/Mask는 N/A
91 | XY Stage | XY Stage positioning Y offset (Global Align) |  |  | 6 | 1.98 | um | Pass | Performance | Check item |  |  | Recipe/분석프로그램 사용, Metro cal 사용, HDM/PSS/PTR/Mask는 N/A
92 | XY Stage | XY Stage positioning X 1σ (Global Align) |  |  | 1.2 | 0.48 | um | Pass | Performance | Check item |  |  | Recipe/분석프로그램 사용, Metro cal 사용, HDM/PSS/PTR/Mask는 N/A
93 | XY Stage | XY Stage positioning Y 1σ (Global Align) |  |  | 1.2 | 0.37 | um | Pass | Performance | Check item |  |  | Recipe/분석프로그램 사용, Metro cal 사용, HDM/PSS/PTR/Mask는 N/A
94 | Z Stage | Z Drift |  |  | 20 | 9.8 | nm | Pass | Performance | Check item |  |  | Zeroscan recipe와 함께 측정되게 변경(v3.13.0), ZeroScan_Zdrift Recipe 사용, TSH사용은 Max spec. 45um 이외는 Max spec. 20um | ☑
95 | ATX(Auto Align Head) | A＋B difference during 3 times run test | -0.25 |  | 0.25 | 0.07 | V | Pass | Performance |  |  |  |  | ☑
96 | ATX(Auto Align Head) | A－B difference during 3 times run test | -1 |  | 1 | 0.5 | V | Pass | Performance |  |  |  |  | ☑
97 | EFEM | Check doubled wafer detection |  | N/A |  | N/A | - | Pass | Performance |  |  |  | Brooks robot인경우 안될수 있음
98 | EFEM | Check crossed wafer detection |  | Yes |  | Yes | - | Pass | Performance |  |  |  | Brooks robot인경우 안될수 있음
99 | EFEM | Wafer placement(Loading) error on the load port1 (X offset) |  |  | 100 | 20.5 | um | Pass | Performance | Check item |  |  | 20번 진행후 Position error 확인 (Max offset 기입), guide 참고
100 | EFEM | Wafer placement(Loading) error on the load port1 (Y offset) |  |  | 100 | 31.1 | um | Pass | Performance | Check item |  |  | 20번 진행후 Position error 확인 (Max offset 기입), guide 참고
101 | EFEM(Dual Load port) | Wafer placement(Loading) error on the load port2 (X offset) |  |  | 100 | 33.9 | um | Pass | Performance | Check item |  |  | 20번 진행후 Position error 확인 (Max offset 기입), guide 참고, loadport2 없을시 삭제
102 | EFEM(Dual Load port) | Wafer placement(Loading) error on the load port2 (Y offset) |  |  | 100 | 15 | um | Pass | Performance | Check item |  |  | 20번 진행후 Position error 확인 (Max offset 기입), guide 참고, loadport2 없을시 삭제
103 | EFEM | Cleanness of EFEM w/ airborne particle measurement |  | ISO1 |  | ISO1 | - | Pass | Performance
104 | EFEM | Angle offset value of pre-aligner |  |  |  | 359.9 | degree | Pass | Performance |  |  |  | EFEM-Alinger 로딩 후 틀어진 각 조정하여 저장한 값 기입(XEA EFEM- Setup - Align Angle offset)
105 | EFEM(Bluecord aligner) | Centering accuracy of pre-aligner |  |  | 0.05 | 0.05 | degree | Pass | Performance
106 | EFEM | Differential pressure measured at FFU Normal speed(Static conditions) | 1.25 |  | 5 | 3.42 | Pa | Pass | Performance |  |  |  | 모든 Panel과 Door 가 닫힌 상태에서 측정, 분석시트 사용
107 | EFEM | Differential pressure measured at FFU Normal speed(Dynamic conditions) | 1.25 |  | 5 | 2.92 | Pa | Pass | Performance |  |  |  | Loadport Door open/close 1회 수행 후 30초간 측정, 분석시트 사용
108 | EFEM | FFU Optimized Velocity Test: Average Airflow Velocity | 0.4 |  | 0.45 | 0.431 | m/s | Pass | Performance |  |  |  | 분석시트 사용
346 | EFEM | FFU Optimized Velocity Test: Individual Velocity Deviation |  | Average Airflow ±20% |  | Pass |  | Pass | Performance |  |  |  | 분석시트 사용하여 pass 시 pass로 선택
110 | EFEM(Dual Load port) | Check 500 cycle of wafer load and unload without error |  | Yes |  | Yes | - | Pass | Performance |  |  |  | YMTC 경우, LoadPort 별 500 times(기본은 LoadPort 별 250times)
111 | Facility Controller(UPS) | Check UPS capability for 5 minutes |  | Yes |  | Yes | - | Pass | Performance
112 | System Noise | Verify MicrometerPerErrorSignalAdc setting for AC160TS |  | 0.2 |  | 0.2 | - | Pass | Calibration |  |  |  |  | ☑
113 | AFM Head | Specification of Step Height Sample for Z Calibration | 458.4 | 461.2 | 464 | 461.2±2.8 | nm | Pass | Calibration |  |  |  | 아래 측정(Z Calibration) 에 사용하는 Sample(450nm) spec 값 min, max criteria에 기입 | ☑
114 | AFM Head | Z Height calibration error |  |  | 0.5 | 0.12576 | % | Pass | Calibration | Check item |  |  |  | ☑
115 | AFM Head | - Z Height measurement value of the calibration sample |  | - |  | 461.78 | nm | Pass | Calibration |  |  |  |  | ☑
116 | AFM Head | Z Drive calibration error |  |  | 0.5 | 0.19731 | % | Pass | Calibration | Check item |  |  |  | ☑
117 | AFM Head | - Z Drive measurement value of the calibration sample |  | - |  | 462.11 | nm | Pass | Calibration |  |  |  |  | ☑
118 | AFM Head(Long) | Z Height calibration value in Large Mode (MicrometerPerAdcHigh in ZDetector DB) | 43 |  | 92 | 66.51 | - | Pass | Calibration | Check item |  | Dsp.ZDetector.LongHead.MicrometerPerAdcHigh |  | ☑
119 | AFM Head(Long) | Z Drive calibration value in Large Mode (MicrometerPerDacHigh in ZScanner DB) | 24 |  | 28 | 26.71 | - | Pass | Calibration | Check item |  |  |  | ☑
120 | AFM Head(Long) | Z Height calibration value in Small Mode (MicrometerPerAdcLow in ZDetector DB) |  |  |  | 66.39 | - | Pass | Calibration | Check item |  |  |  | ☑
121 | AFM Head(Long) | Z Drive calibration value in Small Mode (MicrometerPerDacLow in ZScanner DB) |  |  |  | 26.9 | - | Pass | Calibration | Check item |  |  |  | ☑
122 | Vision | Field Of View in large view | 800 |  |  | 839.6 | - | Pass | Calibration |  |  |  | FieldOfView in Vision Carmera DB | ☑
123 | XY Scanner | X Detector stroke calibration error (nominal scale error X) |  |  | 0.5 | 0.07 | % | Pass | Calibration | Check item |  |  | 신규 분석 프로그램인 XY Scanner Calibration Sample Analyzer 사용 | ☑
124 | XY Scanner | Y Detector stroke calibration error (nominal scale error Y) |  |  | 0.5 | 0.42 | % | Pass | Calibration | Check item |  |  | 신규 분석 프로그램인 XY Scanner Calibration Sample Analyzer 사용 | ☑
125 | XY Scanner | X1 Detector calibration value(MicrometerPerAdc in XDetector DB) | 200 |  | 300 | 250.4 | - | Pass | Calibration | Check item |  |  |  | ☑
126 | XY Scanner | X2 Detector calibration value(AuxMicrometerPerAdc in XDetector DB) | 200 |  | 300 | 258.8 | - | Pass | Calibration | Check item |  |  |  | ☑
127 | XY Scanner | Y1 Detector calibration value(MicrometerPerAdc in YDetector DB) | 200 |  | 300 | 257.8 | - | Pass | Calibration | Check item |  |  |  | ☑
128 | XY Scanner | Y2 Detector calibration value(AuxMicrometerPerAdc in YDetector DB) | 200 |  | 300 | 242.2 | - | Pass | Calibration | Check item |  |  |  | ☑
129 | XY Scanner | X Scanner calibration value(MicrometerPerDacHigh in XScanner DB) | 150 |  | 220 | 164.6 | um | Pass | Calibration | Check item |  |  |  | ☑
130 | XY Scanner | Y Scanner calibration value(MicrometerPerDacHigh in YScanner DB) | 150 |  | 220 | 165.4 | um | Pass | Calibration | Check item |  |  |  | ☑
131 | XY Scanner | Closed loop scanning range (ServoScanRangeMicrometer in X/YScanner DB) |  | 101.01 |  | 101.01 | um | Pass | Calibration |  |  |  |  | ☑
132 | XY Stage | Align the XY stage origin and sample origin |  | Yes |  | Yes | - | Pass | Calibration
133 | XY Stage | RMS error after Stage Calibration |  |  | 5 | 1.8 | um | Pass | Calibration | Check item |  |  | 2번 진행, 2번째 값 기입, N/A시 MAX 값 삭제
134 | Acoustic Enclosure | Check the specification of the nameplate (Refer to the drawing of layout and facility) |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | 제원명판 제작 및 부착
135 | Acoustic Enclosure | AE type |  | Double Walled |  | Double Walled | - | Pass | Configuration |  |  |  | 분리형 | ☑
136 | Computer | Ethernet card type |  | Non-PoE |  | Non-PoE | - | Pass | Configuration |  |  |  |  | ☑
137 | Computer | Disable power saving mode in OS and BIOS |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
138 | Computer | Enable RAID |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
139 | Computer | RAID monitor software |  | Installed |  | Installed | - | Pass | Configuration |  |  |  |  | ☑
140 | Computer | Anti-virus software |  | Installed |  | Installed | - | Pass | Configuration
141 | Computer | Check for the latest update on anti-virus software |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | anti virus software가 최신 업데이트 상태여야함
142 | Computer | Check for the number of viruses detected after the virus scan |  | 0 |  | 0 | - | Pass | Configuration |  |  |  | scan된 virus 개수 확인, virus가 0개여야 함
143 | Computer | Dongle for pattern recognition |  | Installed |  | Installed | - | Pass | Configuration |  |  |  |  | ☑
144 | Computer | Windows version |  | 24H2 |  | 24H2 | Version | Pass | Configuration |  |  |  | Window OS 선택하면 Version 알아서 바뀜, window10 -> 22h2, window11 -> 24h2 | ☑
145 | Computer | Disable Windows update |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
146 | Computer | Save computer configuration file |  | Yes |  | Yes | - | Pass | Configuration
147 | Computer | Check ethernet configuration for AFM Controller |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | Address:192.168.0.1, Subnet mask: 255.255.255.0 | ☑
148 | Computer | Check ethernet configuration for on-axis camera |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | Address:169.254.100.1, Subnet mask: 255.255.255.0 | ☑
149 | Computer | OS(Operating System) type |  | Window11 64bit |  | Window11 64bit | - | Pass | Configuration |  |  |  |  | ☑
150 | Computer | OS / MS-office language |  | English |  | English | - | Pass | Configuration |  |  |  |  | ☑
151 | Computer | Check DotNET Framework Version | 4.6 |  |  | 4.8.1 | Version | Pass | Configuration |  |  |  | Version 4.6 이상인지 확인, ([PST] Environmental Check 프로그램 사용하여 확인) | ☑
152 | Computer | Check the installation status of the display refresh program |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | Display refresh 프로그램 설치 확인 | ☑
153 | Computer | Computer model |  | HP Z4 G5 Workstation |  | HP Z4 G5 Workstation | - | Pass | Configuration |  |  |  |  | ☑
154 | Computer | Check that the right-angle adapter for the HDMI cable is applied |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | HDM/PSS/PTR은 N/A | ☑
155 | Computer | Serial Multi port type |  | Slot |  | Slot |  | Pass | Configuration |  |  |  | 선택하면 아래 Bracket 여/부 바뀜 | ☑
156 | Computer | Check if USB serial multi-port bracket is installed |  | N/A |  | N/A |  | Pass | Configuration |  |  |  |  | ☑
157 | Misc | Check if the mouse and keyboard are functioning properly after turning off the power and then turning it back on |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | 전체 전원 Off 후 10분뒤 전원 On - 마우스,키보드 등 동작 이상없는지 확인 | ☑
158 | Misc | Check that all ground cables are equipped with eye caps |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
159 | Misc | Check that all ground cable labels indicate “PE” |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
160 | Misc | Activate Thermal tune option in XEA |  | Activated |  | Activated | - | Pass | Configuration |  |  |  | Guide참고, xea.vmoption 활성화 및 XEA setup에서 use설정 | ☑
161 | Misc | Resonance frequency of AC160 by thermal tune |  | - |  | 302.51 | kHz | Pass | Configuration |  |  |  | Guide참고, Reference tip을 사용해서 thermal tune과 NCM Sweep capture 저장, Thermal Tune 측정 Data 분석시트에 excel 화 필요(noise 분석용) | ☑
162 | Misc | Resonance frequency of AC160 by frequency sweep |  | - |  | 302.62 | kHz | Pass | Configuration |  |  |  | Guide참고, Reference tip을 사용해서 thermal tune과 NCM Sweep capture 저장, Thermal Tune 측정 Data 분석시트에 excel 화 필요(noise 분석용) | ☑
163 | Misc | Resonance frequency of AC240 by thermal tune |  | - |  | 86.11 | kHz | Pass | Configuration |  |  |  | Guide참고, Reference tip을 사용해서 thermal tune과 NCM Sweep capture 저장, Thermal Tune 측정 Data 분석시트에 excel 화 필요(noise 분석용) | ☑
164 | Misc | Resonance frequency of AC240 by frequency sweep |  | - |  | 86.21 | kHz | Pass | Configuration |  |  |  | Guide참고, Reference tip을 사용해서 thermal tune과 NCM Sweep capture 저장, Thermal Tune 측정 Data 분석시트에 excel 화 필요(noise 분석용) | ☑
165 | Misc | Delete unnecessary files and usage records from the computer according to internal guidelines |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | PC 정리 가이드 참고
166 | Misc | Check SW & DB backup (Installed Software, XEA & XEService) |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | CIM에 Software 및 DB Backup
167 | Misc | Skirt |  | Installed |  | Installed | - | Pass | Configuration
168 | Misc | Seismic bracket |  | Installed |  | Installed | - | Pass | Configuration
169 | Misc | Check if the bolts used for the AFM core module are assembled with standardized torque |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | 옆에 표 참고 | ☑
170 | Misc | Check if the cable holders are installed on both the AE and the Cabinet |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
171 | Misc | Check the AE FFU cleaning |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | AE FFU가 없는 경우는 N/A
172 | Misc | Check the jig bolts for the computer |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
173 | Misc | Check the fixation of SECS/GEM cable on the rear of the PC |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
174 | Misc | Check that all cables connected to the AFM Controller are securely connected with no looseness or detachment |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
175 | Misc | Check that torque markings are applied to the designated fastening bolts |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | 모든 체결볼트 I-marking 확인
176 | Misc | Check that Loctite has been applied to the designated fastening bolts on moving components |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | 모든 구동부 체결볼트 Locktite 적용확인 | ☑
177 | Misc | Check the crimping condition and tension of connectors |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | 모든 커넥터 압착상태 및 장력 확인
178 | Misc | Check that the power supply fastening check sheet is stored in the equipment |  | N/A |  | N/A |  | Pass | Configuration |  |  |  | 삼성향 장비 전원 공급부(Cabinet)에 체결체크시트 격납 확인 이외 장비는 N/A
179 | Misc | Check that temperature and humidity sensor monitoring is enabled and logs are being saved |  | N/A |  | N/A |  | Pass | Configuration |  |  |  | XEA 온습도 모니터링 활성화 및 Log cim 저장
180 | Misc(T1, T3) | Measured resistance value between mounting bolt of Facility Controller and A/E holder ground bar |  |  | 0.5 | 0.2 | Ω | Pass | Configuration |  |  |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인 | ☑
181 | Misc(T1, T3) | Measured resistance value between Z Scanner body and A/E Cable holder |  |  | 1 | 0.5 | Ω | Pass | Configuration |  |  |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인 | ☑
182 | Misc(T1, T3) | Measured resistance value between XY Scanner and A/E Cable holder |  |  | 0.5 | 0.3 | Ω | Pass | Configuration |  |  |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인 | ☑
183 | Misc(T1, T3) | Measured resistance value between ATX Top Plate and A/E Cable holder |  |  | 0.5 | 0.4 | Ω | Pass | Configuration |  |  |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인 | ☑
184 | Misc(T1, T3) | Measured resistance value between XY Stage and A/E Cable holderer |  |  | 0.5 | 0.3 | Ω | Pass | Configuration |  |  |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인 | ☑
185 | Misc(T1, T3) | Measured resistance value between Focus Stage and A/E Cable holder |  |  | 1 | 0.3 | Ω | Pass | Configuration |  |  |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인 | ☑
186 | Misc(T1, T3) | Measured resistance value between Z Stage and A/E Cable holder |  |  | 0.5 | 0.1 | Ω | Pass | Configuration |  |  |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인 | ☑
187 | Misc(T1, T3) | Measured resistance value between Frame Board and A/E Cable holder |  |  | 0.5 | 0.3 | Ω | Pass | Configuration |  |  |  | 전원 off 상태에서 Ground 확인/ NX-Wafer ground map Guide 확인 | ☑
188 | Misc(VLSI) | Check the VLSI sample model number and serial number consistency across the box, certification, and sample |  | - |  | SHS8-9400/30775-008 | - | Pass | Configuration |  |  |  | 고객에게 나가는 모든 샘플(VLSI)에 대한 모델넘버/시리얼넘버 확인하여 measurement에 기입 및 사진으로 남겨 서버에 저장(box, certifiation, sample) | 고객에게 나가는 모든 샘플(VLSI)에 대한 모델넘버/시리얼넘버 확인하여 measurement에 기입 및 사진으로 남겨 서버에 저장(box, certifiation, sample)
189 | Misc(VLSI) | Check the VLSI sample model number and serial number consistency across the box, certification, and sample |  | - |  | SHS8-4500/30812-010 | - | Pass | Configuration |  |  |  | 고객에게 나가는 모든 샘플(VLSI)에 대한 모델넘버/시리얼넘버 확인하여 measurement에 기입 및 사진으로 남겨 서버에 저장(box, certifiation, sample) | 고객에게 나가는 모든 샘플(VLSI)에 대한 모델넘버/시리얼넘버 확인하여 measurement에 기입 및 사진으로 남겨 서버에 저장(box, certifiation, sample)
190 | Misc(w/ EFEM) | Check that there is no disconnection when using the mouse and keyboard inside the EFEM |  | Yes |  | Yes | - | Pass | Configuration
191 | SW | 'Xms4096m' in xea.vmoptions |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | Park systems - XEA - Bin - xea.vmoptions | ☑
192 | SW | 'Xmx16384m' in xea.vmoptions |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | Park systems - XEA - Bin - xea.vmoptions | ☑
193 | SW | 'Xms128m' in xei.vmoptions |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | Park systems - XEI - Bin - xei.vmoptions | ☑
194 | SW | 'Xmx2048m' in xei.vmoptions |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | Park systems - XEI - Bin - xei.vmoptions | ☑
195 | SW | 'Dcom.psia.xea.useInDieAlign=true' in xea.vmoptions |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | Park systems - XEA - Bin - xea.vmoptions | ☑
196 | SW | JAVA JRE version |  | jre-8u191 |  | jre-8u191 | Version | Pass | Configuration |  |  |  |  | ☑
197 | SW | Head Mode in XEA |  | C-AFM, NC-AFM |  | C-AFM, NC-AFM | - | Pass | Configuration |  |  |  |  | ☑
198 | SW | SmartAnalysis version |  | 1.4.3.1480 |  | 1.4.3.1480 | Version | Pass | Configuration |  |  |  | 최신 RTM 버전 사용 | ☑
199 | SW | Z scale in XEA |  | 0.7 |  | 0.7 | - | Pass | Configuration |  |  |  |  | ☑
200 | SW | XEA version |  | 5.2.6.RTM4 |  | 5.2.6.RTM4 | Version | Pass | Configuration |  |  |  | 5.2.4 빌드버전, YMTC 제외 | ☑
201 | SW | XEService version |  | 5.2.6.RTM4 |  | 5.2.6.RTM4 | Version | Pass | Configuration |  |  |  | 5.2.4 빌드버전, YMTC 제외 | ☑
202 | SW | XEI version |  | 5.2.4 B105 |  | 5.2.4 B105 | Version | Pass | Configuration |  |  |  |  | ☑
203 | SW | XEPi version |  | 5.2.0 B28 |  | 5.2.0 B28 | Version | Pass | Configuration |  |  |  |  | ☑
204 | SW | Cantilever DB version |  | 1.3 |  | 1.3 | Version | Pass | Configuration |  |  |  |  | ☑
205 | SW | Head Crash DB setting: A+B Interlock |  | 0.5 |  | 0.5 | V | Pass | Configuration |  |  |  | DB: DSP>Cantilever>MinimalIntensity (Cantiliver: AC160TS) | ☑
206 | SW | Head Crash DB setting: A-B Interlock |  | 6 |  | 6 | V | Pass | Configuration |  |  |  | DB: DSP>Cantilever>SafeErrorSignal (Cantiliver: AC160TS) | ☑
207 | SW | Head Crash DB setting: Z Detector Interlock |  | 0.15 |  | 0.15 | - | Pass | Configuration |  |  |  | DB: DSP/ZStage/HeadCrashThreshold | ☑
208 | SW | Head Crash DB setting: XY Stage Interlock |  | 1 |  | 1 | - | Pass | Configuration |  |  |  | DB: XYStage/Info/bUseInterlock | ☑
209 | System Noise | # of AVIS Element |  | 6 |  | 6 | ea | Pass | Configuration |  |  |  |  | ☑
210 | System Noise | Lateral(C-D) sensitivity |  |  |  | 16.6 | mV/nm | Pass | Configuration | Check item |  |  | B100-NCH A-B Calibration(F/D Curve)
211 | AFM Controller | AFM Controller type |  | IND-M1 |  | IND-M1 | - | Pass | Configuration |  |  |  |  | ☑
212 | AFM Controller | Firmware version of Motor board |  | 0x6e310420 |  | 0x6e310420 | Version | Pass | Configuration |  |  |  |  | ☑
213 | AFM Controller | Firmware version of Base board |  | 0x6e110420 |  | 0x6e110420 | Version | Pass | Configuration |  |  |  |  | ☑
214 | AFM Controller | Line voltage selection |  | 230 |  | 230 | VAC | Pass | Configuration |  |  |  |  | ☑
215 | AFM Head | Check the warranty sticker |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
216 | AFM Head | Z Detector offset |  |  | 200 | 25.7 | nm | Pass | Configuration |  |  |  |  | ☑
217 | AFM Head | AFM Head type |  | Auto Long |  | Auto Long | - | Pass | Configuration |  |  |  |  | ☑
218 | AFM Head | Probehand type |  | High Frequency |  | High Frequency | - | Pass | Configuration |  |  |  |  | ☑
219 | AFM Head | NCMDriveRangeLiquidMode |  | 1 |  | 1 | - | Pass | Configuration |  |  |  | Dsp/Info/NCMDriveRangeLiquidMode=1 | ☑
220 | AFM Head | bDisableNCMDriveReduction |  | 1 |  | 1 | - | Pass | Configuration |  |  |  | Dsp/Info/bDisableNCMDriveReduction=1 | ☑
221 | AFM Head | Check the material of the TSH laser align knob |  | N/A |  | N/A | - | Pass | Configuration |  |  |  | 노브 재질 확인 | ☑
222 | Focus Stage | Check the focus stage reset |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | SW Limit 설정 후 Focus 맨위/맨아래 쪽에서 Reset 진행시 정상동작 여부 확인 | ☑
223 | Focus Stage | Check for interface between the focus stage and head cover when lowering the focus stage |  | N/A |  | N/A | - | Pass | Configuration |  |  |  | TSH가 아닌 경우는 N/A | ☑
224 | Focus Stage | Check the torque marking on the encoder bracket mounting bolts |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
225 | Focus Stage | Check encoder surface is clean |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
226 | Focus Stage | Encoder model |  | V2BCH |  | V2BCH | - | Pass | Configuration |  |  |  | encoder 육안으로 확인(품명적혀있음), 확인후 기입 | ☑
227 | Focus Stage | Check upper limit sensor function |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
228 | Focus Stage | Check lower limit sensor function |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
229 | Focus Stage | LowerLimit |  | H/W lower limit+500um |  | 1624.75 | um | Pass | Configuration |  |  |  | HW limit+500um | ☑
230 | Focus Stage | Check encoder reader LED |  | Blue |  | Blue | - | Pass | Configuration |  |  |  | V2BCH : Good-Blue | ☑
231 | Focus Stage | UpperLimit |  | 9000 |  | 9000 | um | Pass | Configuration |  |  |  |  | ☑
232 | Focus Stage | AutoFocus Range (Tip Focus / Sample Focus) |  | 400 |  | 400 | um | Pass | Configuration |  |  |  |  | ☑
233 | Focus Stage | AutoFocus Speed (Tip Focus / Sample Focus) |  | 150 |  | 150 | um/s | Pass | Configuration |  |  |  |  | ☑
234 | Focus Stage | Fine Autofocus Range (Tip Focus / Sample Focus) |  | 50 |  | 50 | um | Pass | Configuration |  |  |  |  | ☑
235 | Focus Stage | Fine Autofocus Speed (Tip Focus / Sample Focus) |  | 30 |  | 30 | um/s | Pass | Configuration |  |  |  |  | ☑
236 | Focus Stage | UseEncoder |  | 1 |  | 1 | - | Pass | Configuration |  |  |  | FocusStage in DB | ☑
237 | Focus Stage | EncoderResolution |  | 0.05 |  | 0.05 | um | Pass | Configuration |  |  |  | Encoder 품명 육안 확인후 검색하여 Resolution 확인 | ☑
238 | Focus Stage | Fine Autofocus Offset (Tip Focus / Sample Focus) |  | 0 |  | 0 | - | Pass | Configuration |  |  |  |  | ☑
239 | Main Body | Temperature sensor on Z stage |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
240 | Main Body | Firmware version of Frame board |  | 0x6e250420 |  | 0x6e250420 | Version | Pass | Configuration |  |  |  |  | ☑
241 | Sample Chuck | Check chuck surface no damage |  | Yes |  | Yes | - | Pass | Configuration
242 | Sample Chuck | Ground connection |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | AL재질 및 eAFM Sample chuck 제외 | ☑
243 | Sample Chuck | Sample chuck type |  | Pin Bar type |  | Pin Bar type | - | Pass | Configuration |  |  |  | Mask는 N/A | ☑
244 | Sample Chuck | Sample chuck material |  | SiC |  | SiC | - | Pass | Configuration |  |  |  |  | ☑
245 | Sample Chuck | Tip position sample in XEA |  | Taught |  | Taught | - | Pass | Configuration |  |  |  |  | ☑
246 | Sample Chuck(Sample holder) | Check sample holder |  | Yes |  | Yes | - | Pass | Configuration
247 | Vision | Check the white balance |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | Tip-Sample Distance가 어느정도 떨어져 있는 상태에서 White Paper에 Focusing -Brightness 80/Contrast 0/Light strength 50 --> White Balance를 여러 번 클릭해서 WB 맞추기(Vision 밝기가 변화 없을 때까지 -> Gray Color) | ☑
248 | Vision | Check that the camera's white balance setting is saved |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | Guide 참고 : Vimba에서 설정진행 | ☑
249 | Vision | Check Camera illumination function |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | Tip이나 Sample focus 후 1. Light Strength100%(Capture) 2. Light strength 0%(Capture) (Contrast 50%, Brightness는 50%) | ☑
250 | Vision | Check sample autofocus function |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
251 | Vision | Check tip auto focus function |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
252 | Vision | Check tip pattern recognition |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
253 | Vision | Check LightSource ON |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
254 | Vision | LightSource(LED) Color |  | White |  | White | - | Pass | Configuration |  |  |  |  | ☑
255 | Vision | Supply power of the on-axis camera through PoE |  | No |  | No | - | Pass | Configuration |  |  |  |  | ☑
256 | Vision | Supply power of the on-axis camera through frame board |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
257 | Vision | Camera model |  | Mako G-507C |  | Mako G-507C | - | Pass | Configuration |  |  |  |  | ☑
258 | Vision | Camera maker |  | AVT_GIGECAM |  | AVT_GIGECAM | - | Pass | Configuration |  |  |  | Carmera in Vision DB | ☑
259 | Vision | Camera alignment |  |  | 5 | 3.2 | um | Pass | Configuration |  |  |  | Vision상에 보여지는 Chip carrier의 상단 축과(비전 상단끝에 걸치도록 조정) 비전 상단 끝부분 영역이 Rotate 없이 수평으로 맞추어져 있는지 Check, 필요시 Camera Angle을 조정 | ☑
260 | Vision | Vision software version |  | 7.0.2 |  | 7.0.2 | Version | Pass | Configuration |  |  |  |  | ☑
261 | Vision | FullPixel |  | 2440 |  | 2440 | - | Pass | Configuration |  |  |  | Carmera in Vision DB | ☑
262 | Vision | FullPixelHeight |  | 1828 |  | 1828 | - | Pass | Configuration |  |  |  | GT2450C : 1830, Mako G-507C : 1828 | ☑
263 | Vision | Camera FPS(FramePerSecond) of large view | 15 |  |  | 25 | FPS | Pass | Configuration |  |  |  | Large view 일때의 FPS 기입 | ☑
264 | XY Scanner | Check the warranty sticker |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
265 | XY Stage | Check the torque marking on the XY stage mounting bolts |  | Yes |  | Yes | - | Pass | Configuration
266 | XY Stage | Check that anti-loosening washers are applied to the XY stage mounting bolts |  | Yes |  | Yes | - | Pass | Configuration
267 | XY Stage | Check that the motor cable of the XY stage does not interfere with the motor |  | Yes |  | Yes | - | Pass | Configuration
268 | XY Stage | Check that the gore cable of the XY stage is equipped with a cable guard |  | N/A |  | N/A | - | Pass | Configuration |  |  |  | 산업용 분리형 wide type XY Stage 적용, 이외는 N/A 처리
269 | XY Stage | Check that a friction-reducing pad is applied to the XY stage cable veyor |  | Yes |  | Yes | - | Pass | Configuration
270 | XY Stage | Check X axis (+) limit sensor function |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
271 | XY Stage | Check X axis (-) limit sensor function |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
272 | XY Stage | Check Y axis (+) limit sensor function |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
273 | XY Stage | Check Y axis (-) limit sensor function |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
274 | XY Stage | Check XY stage reset |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | SW Limit 설정 후 XY Stage 상/하/좌/우 끝쪽에서 Reset 진행시 정상동작 여부 확인
275 | XY Stage | Check X axis encoder surface is clean |  | Yes |  | Yes | - | Pass | Configuration
276 | XY Stage | Check Y axis encoder surface is clean |  | Yes |  | Yes | - | Pass | Configuration
277 | XY Stage | Save XY stage setup file |  | Yes |  | Yes | - | Pass | Configuration
278 | XY Stage | Suction |  | Enabled |  | Enabled | - | Pass | Configuration
279 | XY Stage | Check the sample align |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | circle align, mask align
280 | XY Stage | Encoder model |  | V2BCH |  | V2BCH | - | Pass | Configuration |  |  |  | XY Stage 상판/하판 사이의 encoder 육안으로 확인(품명적혀있음)
281 | XY Stage | Maker of the motion control board |  | HCMNEW |  | HCMNEW | - | Pass | Configuration |  |  |  | Info in XY Stage DB
282 | XY Stage | Check X axis encoder reader LED |  | Blue |  | Blue | - | Pass | Configuration |  |  |  | V2BCH : Good-Blue
283 | XY Stage | Check Y axis encoder reader LED |  | Blue |  | Blue | - | Pass | Configuration |  |  |  | V2BCH : Good-Blue
284 | XY Stage | Encoder Resolution |  | 0.05 |  | 0.05 | um | Pass | Configuration |  |  |  | Encoder 품명 육안 확인후 검색하여 Resolution 확인
285 | XY Stage | BaudRate |  | 38400 |  | 38400 | - | Pass | Configuration |  |  |  | XY Stage controller 없는 경우 N/A
286 | XY Stage | Serial Port # |  | 1 |  | 1 | - | Pass | Configuration |  |  |  | XY Stage controller 없는 경우 N/A
287 | XY Stage(300mmx200mm) | Sample Loading Position to X |  | - |  | -127021 | - | Pass | Configuration |  |  |  | DB값 기입, XY Stage 체결위치 check
288 | XY Stage(300mmx200mm) | Sample Loading Position to Y |  | - |  | 86462.4 | - | Pass | Configuration |  |  |  | DB값 기입, XY Stage 체결위치 check
289 | XY Stage(300mmx200mm) | StrokeX | 300000 |  |  | 306846 | um | Pass | Configuration |  |  |  | StrokeXMax와 StrokeYMin 기입하면 자동 | ☑
290 | XY Stage(300mmx200mm) | StrokeXMax | 150000 |  |  | 153553 | um | Pass | Configuration |  |  |  |  | ☑
291 | XY Stage(300mmx200mm) | StrokeXMin |  |  | -150000 | -153293 | um | Pass | Configuration |  |  |  |  | ☑
292 | XY Stage(300mmx200mm) | StrokeY | 200000 |  |  | 206631 | um | Pass | Configuration |  |  |  | StrokeXMax와 StrokeYMin 기입하면 자동 | ☑
293 | XY Stage(300mmx200mm) | StrokeYMax | 100000 |  |  | 103162 | um | Pass | Configuration |  |  |  |  | ☑
294 | XY Stage(300mmx200mm) | StrokeYMin |  |  | -100000 | -103469 | um | Pass | Configuration |  |  |  |  | ☑
295 | XY Stage(Sample: 200mm) | StageSafetyX | 200000 |  |  | 200490 | um | Pass | Configuration |  |  |  |  | ☑
296 | XY Stage(Sample: 200mm) | StageSafetyXMax | 100000 |  | 105000 | 100227.8 | um | Pass | Configuration |  |  |  |  | ☑
297 | XY Stage(Sample: 200mm) | StageSafetyXMin | -105000 |  | -100000 | -100190.7 | um | Pass | Configuration |  |  |  |  | ☑
298 | XY Stage(Sample: 200mm) | StageSafetyY | 200000 |  |  | 201304.1 | um | Pass | Configuration |  |  |  |  | ☑
299 | XY Stage(Sample: 200mm) | StageSafetyYMax | 100000 |  | 105000 | 100664.3 | um | Pass | Configuration |  |  |  |  | ☑
300 | XY Stage(Sample: 200mm) | StageSafetyYMin | -105000 |  | -100000 | -100585.9 | um | Pass | Configuration |  |  |  |  | ☑
301 | Z Stage | Check upper limit sensor function |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
302 | Z Stage | Check lower limit sensor function |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
303 | Z Stage | Check the Z stage reset |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | SW Limit 설정 후 Z Stage 맨위/맨아래 쪽에서 Reset 진행시 정상동작 여부 확인 | ☑
304 | Z Stage | Check encoder surface is clean |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
305 | Z Stage | Encoder model |  | V2BCH |  | V2BCH | - | Pass | Configuration |  |  |  | encoder 육안으로 확인(품명적혀있음) | ☑
306 | Z Stage | Check encoder reader LED |  | Blue |  | Blue | - | Pass | Configuration |  |  |  | V2BCH : Good-Blue | ☑
307 | Z Stage | Check the torque marking on the encoder bracket mounting bolts |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
308 | Z Stage | LowerLimit |  | 1Port, 1Slot Z Stage Position -300um |  | 12532 | um | Pass | Configuration |  |  |  | 1Port 1Slot -300um, ATX 없는 경우 Samplechuck approach position -500um | ☑
309 | Z Stage | EncoderResolution |  | 0.05 |  | 0.05 | um | Pass | Configuration |  |  |  | Encoder 품명 육안 확인후 검색하여 Resolution 확인 | ☑
310 | Z Stage | UpperLimit |  | 29000 |  | 29000 | um | Pass | Configuration |  |  |  |  | ☑
311 | Z Stage | UseEncoder |  | 1 |  | 1 | - | Pass | Configuration |  |  |  | ZStage in DB | ☑
312 | Z Stage | HeadCrashThreshold |  | 0.15 |  | 0.15 | - | Pass | Configuration |  |  |  | ZStage in DB | ☑
313 | Z Stage | Initial Z | 12000 |  | 29000 | 20000 | um | Pass | Configuration |  |  |  | Panel: min22000, max29000 | ☑
314 | Z Stage | Z Stage material |  | S45C |  | S45C | - | Pass | Configuration |  |  |  | 옆 사진 참고하여 해당 두가지 파트에 자석이 각각 붙는지 확인한다. (붙으면 s45c, 아니면 Fail) | ☑
315 | Z Stage | Check for interference between the Z and F encoder brackets at the Z stage hardware limit position. |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
316 | ATX | Check if the balls are attaching in each slot |  | Yes |  | Yes | - | Pass | Configuration
317 | ATX | Check if the magnet can move up and down using a software control |  | Yes |  | Yes | - | Pass | Configuration
318 | ATX | Check for CDA leakage |  | Yes |  | Yes | - | Pass | Configuration
319 | ATX | Check teaching position of all slots |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
320 | ATX | Check Tip checker position |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
321 | ATX | Check Probe hand checker |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
322 | ATX | Check ATX cover operation & clearance |  | N/A |  | N/A |  | Pass | Configuration |  |  |  | Cover opens/closes smoothly. No physical interference (rubbing, catching) with other components or chassis, PTR 외에는 N/A | ☑
323 | EFEM | Check if the SID on the EFEM QA report matches the SID on the equipment nameplate |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | EFEM 업체 QA 리포트의 SID와 실제 장비의 명판 SID와 같은지 확인
324 | EFEM | Check if there is marking for robot initial position |  | Yes |  | Yes | - | Pass | Configuration
325 | EFEM | Check the safety interlock |  | Yes |  | Yes | - | Pass | Configuration
326 | EFEM | Check the load port indicator lamp |  | Yes |  | Yes | - | Pass | Configuration
327 | EFEM | Check clearance of robot and end effector against other module |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | Loading 시 간섭여부 Check
328 | EFEM | Check the cassette ID reader |  | Yes |  | Yes | - | Pass | Configuration
329 | EFEM | Check the slot mapping |  | Yes |  | Yes | - | Pass | Configuration
330 | EFEM | Check load and unload from all slot |  | Yes |  | Yes | - | Pass | Configuration
331 | EFEM | Torque of fixing screws for End-Effector |  | 7 |  | 7 | kgf·cm | Pass | Configuration |  |  |  | Robot 및 End effector별 토크기준 상이, Yaskawa Robot : 5kgf·cm/Tazmo : 7kgf·cm
332 | EFEM | Check the light curtain detection |  | Yes |  | Yes | - | Pass | Configuration
333 | EFEM | Check that a cap is installed on the EFEM door lock bolt |  | Yes |  | Yes | - | Pass | Configuration
334 | EFEM | Check that a differential pressure control shutter is applied to the EFEM |  | N/A |  | N/A | - | Pass | Configuration |  |  |  | 300mm Wafer 장비인 경우 Shutter 필요, 그외는 N/A
335 | EFEM | Check that an ESD film is attached to the door window of the EFEM |  | N/A |  | N/A | - | Pass | Configuration |  |  |  | Intel 향인 경우 EFEM Door에 ESD 필름 부착되어있는지 확인, 그외는 N/A
336 | EFEM | EFEM CDA pressure | 0.4 |  | 0.6 | 0.48 | Mpa | Pass | Configuration
337 | EFEM | EFEM vacuum pressure | -90 |  | -60 | -74 | kpa | Pass | Configuration
338 | EFEM | Indexer SW version |  |  |  | 5.2.4 Build60 | Version | Pass | Configuration |  |  |  | KT : Index 5.2.4 Build60 (EFME Robot Disconnect issue 개선 SW)
339 | EFEM | Check Aligner changes angle reference according to software configuration |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | 0deg,90deg,180deg,270deg
340 | EFEM | Check enable pre-load |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | 출하전, XEA에서 Test
341 | EFEM | Level to X axis on the load port 1 |  |  |  | 0.132 | degree | Pass | Configuration |  |  |  | 측정값기입
342 | EFEM | Level to Y axis on the load port 1 |  |  |  | -0.035 | degree | Pass | Configuration |  |  |  | 측정값기입
343 | EFEM | Level to X axis on the robot body | -0.1 |  | 0.1 | -0.01 | degree | Pass | Configuration |  |  |  | 검사성적서 값 기입, 로체(Super plus) EFEM인 경우 Min, Max 값 삭제 후 Criteria에 Center로 표시(Analog leveler bubble 위치임)
344 | EFEM | Level to Y axis on the robot body | -0.1 |  | 0.1 | 0.01 | degree | Pass | Configuration |  |  |  | 검사성적서 값 기입, 로체(Super plus) EFEM인 경우 Min, Max 값 삭제 후 Criteria에 Center로 표시(Analog leveler bubble 위치임)
345 | EFEM | Level to X axis on the robot Upper arm | -0.15 |  | 0.15 | 0.09 | degree | Pass | Configuration |  |  |  | 검사성적서 값 기입, 로체(Super plus) EFEM인 경우 Min, Max 값 삭제 후 Criteria에 Center로 표시(Analog leveler bubble 위치임), Yaskawa type/Non-Yaskawa type spec 다름(메모참고)
346 | EFEM | Level to Y axis on the robot Upper arm | -0.15 |  | 0.15 | -0.06 | degree | Pass | Configuration |  |  |  | 검사성적서 값 기입, 로체(Super plus) EFEM인 경우 Min, Max 값 삭제 후 Criteria에 Center로 표시(Analog leveler bubble 위치임), Yaskawa type/Non-Yaskawa type spec 다름(메모참고)
347 | EFEM | Measurement of surface resistance of the upper end-effector |  | N/A |  | N/A | Ω | Pass | Configuration |  |  |  | Pad type End-effector 만 PAD 표면저항 측정(Trek 사의 표면저항측정기(Model 152-1)을 사용하여 100V 전압을 가해 측정), 아닌경우 N/A 선택하고 Min/Max 값 삭제
348 | EFEM | Measurement of surface resistance of the lower end-effector |  | N/A |  | N/A | Ω | Pass | Configuration |  |  |  | Pad type End-effector 만 PAD 표면저항 측정(Trek 사의 표면저항측정기(Model 152-1)을 사용하여 100V 전압을 가해 측정), 아닌경우 N/A 선택하고 Min/Max 값 삭제
347 | EFEM | Serial port number for the load port1 |  | 5 |  | 5 | - | Pass | Configuration
350 | EFEM(Dual Load port) | Serial port number for the load port2 |  | 6 |  | 6 | - | Pass | Configuration
351 | EFEM(Dual Load port) | Level to X axis on the load port 2 |  |  |  | 0.144 | degree | Pass | Configuration |  |  |  | 측정값기입
352 | EFEM(Dual Load port) | Level to Y axis on the load port 2 |  |  |  | 0.049 | degree | Pass | Configuration |  |  |  | 측정값기입
353 | EFEM(BCR) | Serial Port # |  | N/A |  | N/A | - | Pass | Configuration |  |  |  | Dual loadport인 경우 2,4
354 | EFEM(BCR) | BCR(Bar Code Reader) teach |  | Installed |  | Installed | - | Pass | Configuration
355 | EFEM(OCR) | OCR(Sample ID Reader) Score | 198 |  |  | 399 | - | Pass | Configuration |  |  |  | ioss인 경우 : 120 , cognex인 경우 : 198
356 | Facility Controller | Check facility power cable connection |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
357 | Facility Controller | Check main Disconnect Switch On/Off function |  | Yes |  | Yes | - | Pass | Configuration
358 | Facility Controller | Check A/E door sensor |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
359 | Facility Controller | Check Ionizer controller on/off Switch on controller panel |  | Yes |  | Yes |  | Pass | Configuration |  |  |  |  | ☑
360 | Facility Controller | Check on and off of Ionizer |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
361 | Facility Controller | Check on and off of Suction |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
362 | Facility Controller | Check CCTV for EFEM and AFM |  | Yes |  | Yes | - | Pass | Configuration
363 | Facility Controller | Check vibration amplifier |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
364 | Facility Controller | Check vibration sensors |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
365 | Facility Controller | Check A-B interlock alarm |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | Z stage 구동시, A-B값이 ±3V초과시 Alarm 발생확인 | ☑
366 | Facility Controller | Check A+B Interlock alarm |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | DB : ApproachSetPointRatio = 0.6 작을때 interlock이 걸리는지 | ☑
367 | Facility Controller | Check sample sensing on the sample chuck by vacuum pressure |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | 없는경우 N/A | ☑
368 | Facility Controller | Check A/E lamp on and off by door sensor |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | Door sensor로 Lamp ON/OFF 확인 | ☑
369 | Facility Controller | Check AVIS controller power on and off |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
370 | Facility Controller | Check AVIS controller enable w/o red LED flashing (While XY stage stopped) |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | Panel은 N/A | ☑
371 | Facility Controller | Check Lamp/X-Ray Display for photo ionizer |  | N/A |  | N/A | - | Pass | Configuration |  |  |  |  | ☑
372 | Facility Controller | Check photo ionizer head toward to sample and AFM Head |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | 장착되지 않는 장비는 N/A | ☑
373 | Facility Controller | No power trip during power reset or usage |  | No |  | No | - | Pass | Configuration
374 | Facility Controller | EMO cover included |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | - | EMO COVER가 함께 출고되는 중국향 장비는 Yes, 아닌경우 No
375 | Facility Controller | Signal Tower Color |  | RGBW |  | RGBW | - | Pass | Configuration |  |  |  | 사용색상 선택, 순서상관 없음 | ☑
376 | Facility Controller | Photoionizer lamp(X-Ray Display) color |  | White |  | White | - | Pass | Configuration |  |  |  |  | ☑
377 | Facility Controller | Ionizer Peak Power |  | 5 |  | 5 | KeV | Pass | Configuration |  |  |  | 사용 Power 선택 | ☑
378 | Facility Controller | BaudRate |  | 9600 |  | 9600 | - | Pass | Configuration |  |  |  |  | ☑
379 | Facility Controller(w/EFEM) | Upper Temperature Limit |  | 80 |  | 80 | - | Pass | Configuration |  |  |  | TSH600은 N/A | ☑
380 | Facility Controller(w/EFEM) | Lower Temperature Limit |  | 0 |  | 0 | - | Pass | Configuration |  |  |  | TSH600은 N/A | ☑
381 | Facility Controller | PLC Version |  |  |  | 1.6 | Version | Pass | Configuration |  |  |  |  | ☑
382 | Facility Controller | Vacuum pressure on sample chuck | -80 |  | -60 | -64.1 | kPa | Pass | Configuration |  |  |  | TSH600 Min-80, Max-40, TSH1518/2326 Min-12, Max-8 | ☑
383 | Facility Controller | FFU Check : XEA turn FFU ON / OFF function |  | Yes |  | Yes | - | Pass | Configuration
384 | Facility Controller | A/E FFU set on 100% and 20% then check RPM change |  | Yes |  | Yes | - | Pass | Configuration
385 | Facility Controller(w/EFEM) | EFEM FFU set on 100% and 20% then check RPM change |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | TSH600은 N/A
386 | Facility Controller | AFM FFU speed (Normal) |  | 60 |  | 60 | % | Pass | Configuration
387 | Facility Controller | AFM FFU speed (Active) |  | 20 |  | 20 | % | Pass | Configuration
388 | Facility Controller(w/EFEM) | EFEM FFU speed (Normal) |  | - |  | 80 | % | Pass | Configuration | Check item |  |  | TSH600은 N/A, EFEM FFU Normal 상태는 ffu 풍속테스트 전에 설정 필요 분석시트 및 가이드 참고
389 | Facility Controller(w/EFEM) | EFEM FFU speed (Active) |  | 20 |  | 20 | % | Pass | Configuration |  |  |  | TSH600은 N/A
390 | Facility Controller | Cabinet version |  |  |  | T1 | Version | Pass | Configuration |  |  |  |  | ☑
391 | Facility Controller | Check the pneumatic controller(box) sensor settings (Nomal state : Green) |  | Yes |  | Yes | - | Pass | Configuration |  |  |  |  | ☑
392 | Facility Controller | Check the regulator vacuum pressure | -80 |  | -60 | -70 | kPa | Pass | Configuration |  |  |  | Cabinet에 있는 Regulator vacuum 압력 확인, 사용하지 않는 경우 min,max값 삭제후 criteria N/A | ☑
393 | Facility Controller | Buzzer DB setting : SignalTower - bReportDoorOpen​ |  | 1 |  | 1 | - | Pass | Configuration |  |  |  | DB 설정 : FacilityController/​SignTower/
394 | Facility Controller | Buzzer DB setting : SignalTower - bReportEmergency |  | 1 |  | 1 | - | Pass | Configuration |  |  |  | DB 설정 : FacilityController/​SignTower/
395 | Facility Controller | Buzzer DB setting : SignalTower - bUseBuzzer |  | 1 |  | 1 | - | Pass | Configuration |  |  |  | DB 설정 : FacilityController/​SignTower/
396 | Facility Controller | Buzzer DB setting : Emergency - bUse |  | 1 |  | 1 | - | Pass | Configuration |  |  |  | DB 설정 : FacilityController/​Emergency/
397 | Facility Controller | Buzzer DB setting : Emergency - bUseResetAlarm​ |  | 1 |  | 1 | - | Pass | Configuration |  |  |  | DB 설정 : FacilityController/​Emergency/
398 | Facility Controller | Buzzer DB setting : Emergency - bUseSystemReset |  | 1 |  | 1 | - | Pass | Configuration |  |  |  | DB 설정 : FacilityController/​Emergency/
399 | Facility Controller | Buzzer DB setting : Emergency - bUseWarningMessage​ |  | 1 |  | 1 | - | Pass | Configuration |  |  |  | DB 설정 : FacilityController/​Emergency/
400 | Facility Controller | Check the buzzer |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | EMO 동작시 Alarm 창과 부저가 울리지는지 확인
401 | Facility Controller(T1,T2) | Check if the power lamps(Power Cable/UPS Power) turn on when the main power is on |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | power on 했을때 power lamp(Power Cable/UPS Power)가 점등 되는지 확인
402 | Facility Controller(T1,T2) | Check the EFEM EMO switches function test |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | EMO Alarm 이미지 저장, EFEM EMO Switch 위치 3군데 (Teaching pendant, EFEM Front, EFEM Monitor 쪽)
403 | Facility Controller(T1,T2,T3) | Check the Cabinet EMO switches function test |  | Yes |  | Yes | - | Pass | Configuration |  |  |  | EMO Alarm 이미지 저장, Cabinet EMO Switch 위치 2군데 (Cabinet 오른쪽, Cabinet 안쪽)
404 | Facility Controller(T1,T2,T3) | Serial Port # |  | 3 |  | 3 | - | Pass | Configuration |  |  |  | T1,T2 : 3, T3:1(단, PTR은 2) | ☑
405 | Facility Controller(T1,T2,T3) | Power lamp color |  | White |  | White | - | Pass | Configuration |  |  |  |  | ☑
406 | Facility Controller(T1,T2,T3) | Swtich box lamp color(Power On = Green, Power Off = White) |  | Green, White |  | Green, White | - | Pass | Configuration |  |  |  |  | ☑
407 | Facility Controller(UPS) | UPS Power lamp color |  | White |  | White | - | Pass | Configuration
408 | Tip Characterizer Sample | Tip Characterizer stage position teach / XEA |  | Taught |  | Taught | - | Pass | Configuration |  |  |  | Guide 참고
409 | Tip Characterizer Sample | Tip Characterizer DB setting : TipCharacterizerSupport |  | 1 |  | 1 | - | Pass | Configuration |  |  |  | Guide 참고
410 | Tip Characterizer Sample | Tip Characterizer auto focus teach / XEA |  | Taught |  | Taught | - | Pass | Configuration |  |  |  | Guide 참고
411 | Tip Characterizer Sample | Tip Characterizer sample width |  | - |  | 201.3 | nm | Pass | Configuration |  |  |  | Guide 참고
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
411
Note on B8: Soon Jang: NX-Mask는 삭제
Note on B9: Soon Jang: NX-Mask는 삭제
Note on B10: Soon Jang: NX-Mask는 삭제
Note on B11: Soon Jang: NX-Mask는 삭제
Note on B12: Soon Jang: NX-Mask는 삭제
Note on B13: Soon Jang: NX-Mask는 삭제
Note on B14: Soon Jang: Panel 장비는 Dynamic test 삭제
Note on G107: Soon Jang: 설정한 값 기입
Note on G108: Soon Jang: 설정한 값 기입
Note on E256: Soon Jang: Yes or No 선택시 아래 check item Criteria 변경됨
Note on D346: Soon Jang: Yaskawa : ±0.1 degree Non-Yaskawa : ±0.15 degree
Note on F346: Soon Jang: Yaskawa : ±0.10 degree Non-Yaskawa : ±0.15 degree
Note on D347: Soon Jang: Yaskawa : ±0.1 degree Non-Yaskawa : ±0.15 degree
Note on F347: Soon Jang: Yaskawa : ±0.10 degree Non-Yaskawa : ±0.15 degree
Note on G389: Soon Jang: 설정한 값 기입

## Last
Product Quality Conformity Certificate
This is to certify the product listed below and associated components have been manufactured, assembled, and tested by Park Systems, in accordance with established manufacturing procedures and ISO9001 and ISO14001 guidelines.
|  | Product Model | [5 empty] | NX-Wafer 200mm
|  | SID(System ID) Number | [5 empty] | C160112-020725
|  | Reference Document | [5 empty] | Industrial Checklist_2026-Q1-Rev.0
|  | Date of Final Test | [5 empty] | 2026-03-13
|  | End User | [5 empty] | Huahong Grace Fab3 #2
|  | Manufacturing Engineer | [5 empty] | JunBeom Park
|  | Production QC Engineer | [5 empty] | Levi. Baek
|  | Team Manager of Production QC | [5 empty] | Jayden Lee
-
|  | Teached |  | Teached |  | Pass
5.2.4 빌드버전, YMTC 제외
5.2.4 빌드버전, YMTC 제외
|  |  |  | 1 |  | Fail
um
Resonance Frequency in DB
Check XY stage reset
Save XY Stage setup file
Suction |  | Enabled |  | Enabled |  | Pass
Check the circle align
Maker of the motion control board(Maker in XY Stage DB)
Check X axis encoder reader LED(Green: Good, Red: No Good)
Check Y axis encoder reader LED(Green: Good, Red: No Good)
Encoder Resolution
Sample Loading Position to X
Sample Loading Position to Y
Image offset after moving from -X to +X | [6 empty] | Plus direction move Drift, X move X Scan(XX) 700sec (-X to +X move, image offset
Image offset after moving from -Y to +Y
Image offset after moving from -X to +Y
Image offset after moving from -Y to +X
Image offset after moving from +X to -X
Image offset after moving from +Y to -Y
Image offset after moving from +X to -Y
Image offset after moving from +Y to -X
X axis, Repeatability @1st position
X axis, Repeatability @ 2nd position
X axis, Repeatability @ 3rd position
X axis, Repeatability @ 4th position
X axis, Accuracy @1st position | [6 empty] | Y axis, Accuracy @1st position
X axis, Accuracy @ 2nd position | [6 empty] | Y axis, Accuracy @ 2nd position
X axis, Accuracy @ 3rd position | [6 empty] | Y axis, Accuracy @ 3rd position
X axis, Accuracy @ 4th position | [6 empty] | Y axis, Accuracy @ 4th position
Y axis, Repeatability @1st position
Y axis, Repeatability @ 2nd position
Y axis, Repeatability @ 3rd position
Y axis, Repeatability @ 4th position
Y axis, Accuracy @1st position
Y axis, Accuracy @ 2nd position
Y axis, Accuracy @ 3rd position
Y axis, Accuracy @ 4th position
Check the Z stage reset
Z Drift (Short Traveling)
Check the Z Scanner Range
Check the Resolution of 1line (Max Data Length in Profile mode) |  |  |  | 524288
Check the Profiling Speed (Max Speed in Profile mode)
Check the Profiling Length (Max Profiling Length in Profile mode)
Check the Max Scan Speed (Max Scan speed in AFM mode)